simpleSTORM:
a self-calibrating
Single molecule
localization
algorithm
Hamprecht Lab, HCI/IWR
Ilia Kats
April 22, 2013

---
Outline
- Introduction: dSTORM
- simpleSTORM: Overview
- simpleSTORM in detail
- ISBI Localization Microscopy Challenge

---
The optical resoltion limit
FWHM = %lambda over {2 cdot NA}
http://upload.wikimedia.org/wikipedia/commons/c/c2/Convolution_Illustrated_eng.png

---
dSTORM
Heilemann, Mike, van de Linde, Sebastian, Schüttpelz, Mark, Kasper, Robert, Seefeldt, Britta, Mukherjee, Anindita, Tinnefeld, Philip, Sauer, Markus: Subdiffraction-resolution fluorescence imaging with conventional fluorescent probes., Angew. Chem. Int. Ed. Engl. 47(33), 6172–6, 2008
[Speaker Notes: Cy5: excitation at 633nm 500 ms cont. excitation: inactivation for 5 s reactivation: irraditation at 488 nm simultaneous irradiation with 488, 633 nm: blinking adjust intensity such that only a subset of fluorophores is excited image: microtubules in COS7]

---
simpleSTORM:
Overview
Stage 1: estimate gain/offset
Stage 2: adjust gain
Stage 3: estimate PSF width
Stage 4: localize molecules

---
Stage 1:
Gain/offset estimation
Hwang, Youngbae, Kim, Jun-Sik, Kweon, In-So: Difference-Based Image Noise Modeling Using Skellam Distribution., IEEE Trans. Pattern Anal. Mach. Intell. 34(7), 1329–1341, 2012
[Speaker Notes: Skellam distribution: difference of Poisson pixel has no true intensity, is result of random process skellam parameter has much clearer linear dependence of intenstiy than variance Risl: ration of ISL to to intensity, only depends on gain ISL: intensity-skellam line build average of all frames, sample points over complete value range fitting: MCD, fit only over subset gain: slope offset: x0 > minVal ? minVal : x0 offset = 0 if minVal < 0]

---
Stages 2-4:
Preprocessing
I' = {I - offset} over gain newline I'' = 2 cdot sqrt{I' + 3 over 8}
transform to Poisson distribution
Anscombe Transform: approx. Gaussian with unit variance
-
=

---
Stage 2:
Gain Adjustment
[Speaker Notes: iterative algorithm use gain from stage 1 as starting value goal: variance of 1 for the intensity distribution histogram 1: variance = ca. 2 histogrm 2: variance = ca. 1]

---
side-note:
Fourier Transform
Smith, Steven W.: The Scientist and Engineer's Guide to Digital Signal Processing, California Technical Publishing, 1997
[Speaker Notes: infinite number of sinusoids required to synthesize aperiodic signals => only DFT used in signal processing]

---
Side-Note:
Fourier Transform
Smith, Steven W.: The Scientist and Engineer's Guide to Digital Signal Processing, California Technical Publishing, 1997
[Speaker Notes: Nyquist frequency: ½ of samples harmonics over Nyquist frequency yield the same time signal as those below => aliasing real/imaginary: only meaningful for complex DFT]

---
Side-Note:
Fourier Transform
magnitude of the complex number
Smith, Steven W.: The Scientist and Engineer's Guide to Digital Signal Processing, California Technical Publishing, 1997
[Speaker Notes: windowed signal: divide original (long) signal into short windows multiply with smoothing filter, average DFTs of windows => smaller effect of noise]

---
Stage 3:
PSF estimation
[Speaker Notes: FT of gaussian gives gaussian std.dev. inversely proportional select 5 random local maxima per frame, average Power Spectra requirement for maximum: thresh = 0.7 * (max - min) + min min, max: minimal and maximal intensity of image repeat for 200 frames fit Gaussian in frequency domain, compute std.dev. in spatial domain center of FT: mean intensity of image (DC component)]

---
Stage 4:
Molecule localization
threshold
filter
[Speaker Notes: matched filter: improves SNR threshold: based on p-value for normal distribution (0-mean, unit var.) select local maxima above threshold interpolate region around maximum get local maximum, use as molecule localization]

---
Result

---
ISBI Localization microscopy Challenge

---
ISBI Localization Microscopy Challenge
http://bigwww.epfl.ch/smlm/challenge/images/lm.007.png